## Configuration

### Boards

This test contemplates a backplane and a single mux board. It can also be adapted for multiple mux boards.

### Connections

Motherboard Arduino is powered and controlled from USB. (Serial port is also possible but the firmware must be reconfigured.)

Electrometer input connected to Mux board IV out, or to backplane IV out, as needed

### Jumpers

#### Motherboard

JP5: supplies +5V power to shift registers, via 12 – 5V linear regulator U2

Other jumpers uninstalled.

#### MUX board

JP2: connects VB_POS

JP1: install if testing IV_OUT on backplane.

Other jumpers uninstalled.

## Functional Tests

These are some quick tests to:

- determine the leakage of the board

- ensure the relay activation works

- ensure the diodes are the right way around

Board leakage in the bias circuit can be tested prior to plugging in relays. This allows isolation of relay leakage and board leakage.

### Bias Relay Activation

#### Coil Voltage

Check for +12V on pin 2 when the channel is selected.

VB should be present on pin 7 (passed by the photomos), if and only if the channel is selected.

#### Bypass

To bypass the relay (bias circuit), jump pins 1-7.

### Sense Relays

Put the sense relays in both states and check that the relays are alternately closed or open.

### PIN Diodes

## IV Mode Tests

For these tests:

- IV out line is connected to the electrometer (Triax connection TBD)

  - Both mux IV out and backplane IV out must be tested at some point

  - The mux output can be used initially, to remove any issues about the backplane connector

- To protect the electrometer, use a series resistor here or on the Vb source

- Component references are to the first channel or first channel group.

### Noise Level

- Assess susceptibility to nearby movements etc which may create dynamic electric fields that induce noise. Also, 60Hz fields are a possibility, which can be partially canceled by the correct electrometer integration timing setting (power-line cycles)

- Assess electrometer reading noise level and reduce it to the level needed (~1pA?) using electrostatic shielding

- This can be with no voltages applied to the circuit

- This ensures the current reading noise level is negligible before we consider any leakage currents

### Sense Line Leakage from +12V

- Nothing is plugged in to the SiPM connector

- All bias relays off

- Sense relay off (IV_ON low)

- Let electrometer reading settle

- Turn +12V on and look at any change in electrometer reading

- Sense relay on (IV_ON high)

- Turn +12V on and look at any change in electrometer reading

### Bias Relay Leakage from VB, no SiPM connections biased, no connector present

Will measure any leakage from VB to sense, independent of the HD connector.

- Nothing is plugged in to the SiPM connector

- All bias relays off, sense relay on

Apply and remove VB and check for any related change in electrometer current

#### Result 2025-02-03

VB was 50V.

We measured this in the previous week and concluded there was no significant leakage.

We re-measured it today and saw ~6pA which goes to 100-500fA when VB is removed.

### Bias Relay Leakage from +12V, no SiPM connections biased

No test is required, since +12V will only be present on a given relay connection, if it is active.

Leakage current is assumed to be negligible compared with VB connection (via the bias resistors).

### Bias Relay Leakage from VB, no SiPM connections biased, all bias lines inactive but shorted

Will measure any leakage from VB to any inactive cathode bias trace.

- Jumper all of the SiPM connections (K-A) on the Tile Interface (with 2mm jumpers). This may also be done in sets of SiPM connections.

- Plug Tile Interface (directly) into SiPM connector

- Activate the sense relay connection through to the electrometer, using the command e.g. s 1 1 (sense relay, channel 1, on).

Apply and remove VB and check for any related change in electrometer current

Further diagnosis:

- If a change with connection / disconnection of Vb is seen here (but not in test 1.3.1), narrow down which channels (jumpers) are responsible for the leakage. The problem is probably found in the relay circuit.

#### Result 2025-02-03

We did this on the first 8 channels, one at a time.

The TIB was plugged in directly, a

Readings are with moving average filter set to 100 samples. A significant amount of reading noise was seen.

We need to track down whether this comes from

- the TIB or adapter board being connected

- wires being connected to the adapter board

- the wires actually being shorted

| Channel | Reading 1 (pA) | Reading 2 (pA) |
| --- | --- | --- |
| 1 | 60 | 75 |
| 2 | 24 |  |
| 3 | 22 |  |
| 4 | 27 |  |
| 5 | 23 |  |
| 6 | 21 |  |
| 7 | 24 |  |
| 8 | 220 (maybe no moving average) | 94 |

#### Result 2025-02-07

With a power supply of 12V, we see a current of 200 pA. When power supply is turned off, the two big relays have 3.6 Vs across them, and we see a current of about 62pA. This is consistent as (200/62)*12=3.7V.

Having the Arduino powered puts a residual voltage on the 12V pin even then the 12V is absent.

On cutting of the relay pin, we see the leakage goes down to 0.05pA, bouncing back and forth between +/-, with bias and power supply on.

On removing the relay and activating bias channel 1, we see a less than a pA current. --So the relays are an issue.

With the series relay bypassed and the shunt relay disconnected, we see <1pA with no bias populated, and nothing plugged in to the HD connector.

Mystery

When we connect the TIB to the HD connector, we see ~50-60pA (with the 100 sample MA(((((moving aaveraverage)MAmoving average)-moving average)-moving average filter). We didn’t see this in 1.2.5.1. Getting more noise because the TIB and adapter are connected is normal since the anode connection isn’t shielded by proximity to the board. But how can plugging in the TIB result in a net positive current?

Where is the reading noise coming from? Is it leaking in from outside? Is it higher frequency? Is it from the Arduino? Unplugging the Arduino and powering down the 12V, we don’t see any dramatic reduction in reading noise.

With the relays removed and Ch 8 SiPM shorted, we don’t see any evidence of Vb-related leakage current. (Still need to test this with the relays in place.)

To be done: Leakage check on channels 1-8 after sonication of boards

Ch 8-  no significant leakage excess seen when ch8 was activated/ deactivated.

### Leakage through photomos

It will be interesting to sample the leakage through the photomos with Vb = 50V, by measuring the voltage present on R11

### Adjacent Bias Line Leakage

- With all SiPM connections jumpered (K-A) on connected Tile Interface

This test chiefly checks for leakage from any one active cathode bias line, into any of the other cathode lines. Note that although the procedure calls for all of the other channels to be jumpered, it may be adequate to only jumper channels adjacent or close to the active channel.

For each of 24 channels:

- Zero state

- Apply VB

- Remove jumper on connector for channel under test

- Jumper adjacent channels on connector

- Activate the channel and measure

Check for any change in electrometer reading that is related to channel or VB activation

Further diagnosis:

- If significant change is found, determine whether it depends on adjacent jumpers being connected (i.e. it’s via other bias lines). If not, it may be directly on the channel itself.

- If it appears to be directly on the channel, repeat with tile interface removed, as this could be the source of the leakage.

## Pulse Mode Tests

### Noise

The only potential noise source on the board is probably the Arduino. Therefore we can assess whether it noticeably increases the noise floor. However, to turn on a pulse output path, the Arduino is needed (unless we try some hack).

Therefore, let’s do the following to start with.

Look at, capture, and compare the RF output noise level with Arduino on and off.

With Arduino on, activate some channels and compare the RF output noise level.

- Channels terminated with short or 50 ohms (at the HD connector)

- Channels left open

### Attenuation and distortion

- With 50 ohm drive

  - Check attenuation versus straight through with 50 ohm termination

- With high impedance drive

- Look at attenuation and distortion

- Start with 50nS pulse, reduce to shorter pulse

### Supplementary

VNA sweep (frequency)

Cable test sweep