# IV-PULSE MUX

Technical Documentation

Contents

1.	Introduction	3

1.1	High Level Theory of Operation	3

1.2	Block Diagram	3

2.	Top Level Component Description	3

2.1	Mux Backplane (MBP)	3

2.1.1	Interfaces	3

2.1.2	Jumper Configurations	4

2.1.3	Arduino Nano Every	5

2.2	Backplane Heat Sink (BHS)	6

2.3	Mux Switching Board (MSB)	6

2.3.1	Interfaces	6

2.3.2	Jumper Configurations	8

2.3.3	Circuit Details	9

2.3.4	GNDD connection	14

2.4	Mux Board Supports	14

2.5	Heat Sink Board (HSB)	14

2.6	Tile Interface Board (TIB)	15

2.7	MSB to HSB Cable	15

2.8	HSB to TIB Cable	15

2.9	SiPM Tile	15

3.	Notes	15

4.	TDL	16

5.	Earlier Text	17

5.1	Technical Description	17

5.1.1	Bias Control Circuit	17

5.1.2	IV Mode	17

5.1.3	Pulse Mode	17

5.1.4	Components	17

5.1.5	Ideas / Notes / Calculations	19

# Introduction

This document describes the IV-Pulse mux, a system developed for testing SiPMs in a cryogenic (LXe) environment.

It combines low-current IV curve sensing with decent RF characteristics for sensing individual SiPM output pulses. The goal is to enable automatic switching among 96 SiPMs mounted on a test tile, enabling individual SiPMs to be selected for IV curve and pulse counting testing, while minimizing cryostat feedthroughs.

## High Level Theory of Operation

Take text from 1 below

Mention the role of the capacitors in the TIB.

## Block Diagram

To be added, with a brief description.

# Top Level Component Description

Elements of the system are described below.

## Mux Backplane (MBP)

The Mux Backplane (MBP) mounts to the backplane heat sink. It incorporates an Arduino Nano Every, to control the four mux switching boards.

### Interfaces

#### Power and TTL-level serial interface

This is the connection to the outside world.

A 9-pin DSUB connector is used, and the same pinout can be maintained through a UHV bulkhead feedthrough. Note that the serial interface is TTL-level, and this is not an RS-232 compatible connector, neither with regard to voltage or pinout.

In addition, for testing, a 6 pin single-inline header with an FTDI-adapter socket compatible pinout is provided.

| Function | DE9 Pin | FTDI pin |
| --- | --- | --- |
| +12V | 5 | n/a |
| +5V (optional) | 9 | 2 |
| GNDA | 1 | 1 |
| DGND |  |  |
| VB+ | 4 |  |
| VB- (this is probably analog ground) |  |  |
| Serial Tx (from Arduino) | 3 | 5 |
| Serial Rx (to Arduino) | 8 | 4 |
| RTS, may be used for Arduino reset | 7 | 6 |
| GNDA | 0 (shell) |  |

Notes on pin functions:

+12V is passed to the MSBs and is used to activate its relays.

+5V can be provided via the DSUB or FTDI connector. For more information on power options, see 2.1.2.1 below

What do we say about GNDD?

#### MSB Interface

For a detailed description of this interface, see 2.3.1.1.

#### IV curve output

Triax or BNC (TBD)

Note that we may also allow direct interconnect between mux boards for IV measurements, in the event that it’s problematic to connect this net to the backplane.

### Jumper Configurations

#### Power

JP1 connects the 5V rail to the FTDI header J9 pin 2.

JP2 connects the 5V rail to the DE9 pin 9.

JP3 connects the 5V rail to the +5V pin on the Arduino Nano Every

JP4 connects the 12V rail to the VIN pin on the Arduino Nano Every

JP5 connects the output of the 78L05 to the 5V rail

The following table shows 5V rail power jumper configurations:

|  | JP1 | JP2 | JP3 | JP4 | JP5 |
| --- | --- | --- | --- | --- | --- |
| 5V rail from FTDI header | IN |  |  |  |  |
| 5V rail from DE9 |  | IN |  |  |  |
| 5V rail from Arduino |  |  | IN | IN |  |
| 5V rail from 78L05 |  |  |  |  | IN |

The following table shows two options for powering the Arduino:

|  | JP1 | JP2 | JP3 | JP4 | JP5 |
| --- | --- | --- | --- | --- | --- |
| Arduino powered from +5V rail |  |  | IN |  |  |
| Arduino powered from +12V on VIN pin |  |  |  | IN |  |

Note that these options may be overly complicated. One setup that works is as follows:

JP3 IN, JP1, 2, 4, 5 OUT.

This relies on the 1A switching converter on the Nano Every board to source the 5V rail. This will result in optimal efficiency but may produce more noise. If noise is an issue, the 78L05 linear regulator can be used instead.

#### External Reset Function

JP6 configures the Arduino Nano Every reset function, which can optionally be activated using a serial port RTS signal.

With no jumper, there is no external reset function.

Connecting JP6 1-2 will configure the Arduino to reset if the TTL-level active-high RTS line is set inactive-low.

Connecting JP6 2-3 will configure the Arduino to reset if the TTL-level active-high RTS signal is set active-high. For this function to be supported, Q1 must be populated.

### Arduino Nano Every

The Arduino Nano Every is connected to the MBP using a dual-inline socket header. It may also be soldered in place, but this is probably less convenient.

The Arduino has a USB port which may be used to program it, while it is either connected to the MBP or disconnected. The USB port can also be used to control the mux, typically for testing purposes, if the firmware is configured to use the USB virtual COM port rather than the TX1 / RX1 port that is connected to the FTDI header and DE9.

The Arduino includes a reset button that can be used if the Arduino is accessible.

The Arduino includes a switching power converter that can run on VIN (connected to the +12V rail using JP4), and which is nominally capable of outputting 1A, which would normally be adequate for powering both the Arduino itself and everything else on the 5V rail of the system. Some care should be taken to ensure that in a vacuum environment, the Arduino module has adequate heat sinking. It should be possible to read the MCU’s temperature by firmware (TBD whether this is working.)

The following sections describe the Nano Every functions.

#### GPIO Pins

To be listed in a table.

#### Bias Relay Control (Shift Registers)

To be detailed.

#### Sense Relay Control

To be detailed.

#### Thermistor Reading

To be detailed.

#### Any others?

To be detailed.

## Backplane Heat Sink (BHS)

The backplane heatsink is used to mount the backplane. This part is mentioned in the event it is needed to achieve better conductive heat transfer. It may not be needed, and the backplane may simply be mounted on standoffs.

## Mux Switching Board (MSB)

The full system supports testing of 96 SiPMs on one tile, using four MSBs, each of which services one 24-SiPM “quad”.

### Interfaces

#### J2: Backplane (MBP) Connector

The MSB has a DC37 male right angle connector on one end, which interfaces to the MBP.

| Pin | Signal Name | Description |
| --- | --- | --- |
| 0 (shield) | GNDA | Analog GND. Note that many pins are connected to GNDA to promote heat transfer from the ground plane. |
| 1 | GNDA | See pin 0 above |
| 2 | GNDA | See pin 0 above |
| 3 | GNDA | See pin 0 above |
| 4 | GNDA | See pin 0 above |
| 5 | PIN_BIAS2 | Bias voltage for first stage PIN diode cathodes & second state PIN diode anodes |
| 6 | SER_IN | Data input to MSB for shift register |
| 7 | GNDA | See pin 0 above |
| 8 | GNDD | Digital ground, return for relay current |
| 9 | GNDA | See pin 0 above |
| 10 | SER_OUT | Data output from MSB shift register. Not surrently |
| 11 |  | Output enable for the shift registers, active low. |
| 12 | GNDA | See pin 0 above |
| 13 | GNDA | See pin 0 above |
| 14 | THERMSEL2 | Thermistor select code, bit 2 |
| 15 | THERMSEL0 | Thermistor select code, bit 0 |
| 16 | GNDA | See pin 0 above |
| 17 | VB |  |
| 18 | GNDA | See pin 0 above |
| 19 | IV_OUT |  |
| 20 | VSENSE |  |
| 21 | GNDA | See pin 0 above |
| 22 | GNDA | See pin 0 above |
| 23 | PIN_BIAS1 | Bias voltage for second state PIN diode cathodes |
| 24 | RCLK |  |
| 25 | GNDA | See pin 0 above |
| 26 | GNDA | See pin 0 above |
| 27 | GNDD |  |
| 28 | +12V |  |
| 29 | SRCLK |  |
| 30 | +5V |  |
| 31 | GNDA | See pin 0 above |
| 32 | ALT RLY VLT | alternate voltage for bias relays |
| 33 | THERMSEL1 | Thermistor select code, bit 1 |
| 34 | THERM_ON |  |
| 35 | IV_ON |  |
| 36 | THERM_MUX |  |
| 37 | GNDA | See pin 0 above |

#### IV OUT (J4) and RF OUT (J1)

Both of these connectors are vertical MCX socket.

The IV OUT connector can be used to connect the IV current signal to the electrometer input, as an alternative to passing it through the MBP. This would be used in two scenarios:

- If noise or leakage performance were found to be degraded by connecting the four MSB’s IV OUT circuits together, or passing them through the MBP.

- If simultaneous readout from multiple MSB’s were desired, to allow parallel testing of one SiPM one each tile quad.

The RF OUT connector connects to the RF path of the SiPM currently activated by the MSB. All four MSBs are intended to be connected to their own pulse counting circuit, via vacuum feedthroughs.

Cable assemblies with right-angle MCX plugs should typically be used inside the cryostat to mate with the board-side MCXs. This inherently adds a degree of freedom in cable routing, by swiveling to a convenient angle.

Note that pin compatible through-hole vertical SMA connector can also be fitted to the board in these two locations, but it may not be as easy to find a matching cable assembly that has a low enough profile to fit between the boards in the in-cryostat assembly, nor will the cable inherently swivel as with the MCX.

#### SiPM Quad Interface

The SiPM Quad interface consists of a Samtec QSE-040-01-L-D-A “High-Speed Ground Plane Socket Strip”. It consists of two inline rows, each row having 40 pins, as well as several ground strips. This mates with a “Terminal” connector which terminates two coaxial ribbon cables, each ribbon comprising 40 coaxial cables, with the coaxial cable shields connected to the socket ground strips.

Note that in this Samtec connector system, “Terminals” mate with “Sockets”.

The connector carries the following signals:

- GNDA (connected to all coax shields, as well as all pins not listed in the table)

- 24 SiPM cathode connections: SIPM_K0 through SIPM_K23

- Two IV return connections (connected together on the MSB)

- Two thermistor connections for remote temperature sensing

| Pin | Signal Name | Pin | Signal Name |
| --- | --- | --- | --- |
| 2 | SIPM_K7 | 42 | IV return |
| 6 | SIPM_K6 | 45 | SIPM_K13 |
| 10 | SIPM_K5 | 49 | SIPM_K14 |
| 14 | SIPM_K4 | 52 | SIPM_K23 |
| 18 | SIPM_K4 | 53 | SIPM_K15 |
| 22 | SIPM_K2 | 56 | SIPM_K22 |
| 26 | SIPM_K1 | 60 | SIPM_K21 |
| 27 | SIPM_K8 | 64 | SIPM_K20 |
| 30 | SIPM_K0 | 68 | SIPM_K19 |
| 31 | SIPM_K9 | 72 | SIPM_K18 |
| 35 | SIPM_K10 | 76 | SIPM_K17 |
| 39 | SIPM_K11 | 77 | THERM4 |
| 40 | IV return | 79 | THERM5 |
| 41 | SIPM_K12 | 80 | SIPM_K16 |

### Jumper Configurations

#### JP1 – IV_OUT to DSUB

When placed, this jumper connects the IV out signal from the mux, via the DSUB connector J, to the Mux Backplane, so that a single, combined IV_OUT signal from the backplane can be connected to the electrometer. If separate IV_OUT signals are connected from the MSBs using J4, JP1 should not be placed.

#### JP2 – VB resistor bypass

This jumper is placed to bypass the series resistor R135, or to connect VB to the bias switching relays, in the event that the resistor is not populated.

#### JP3 – JP5

These are normally-closed solder jumpers bypassing unused series resistors on U1, U2 and U3 VCC_LOAD supply pins. Nothing needs to be done to them.

#### JP6 – PIN Diode Current Sink

3-way solder jumper, to reconfigure, pin 1-2 must be cut and pins 2-3 bridged.

1-2: sink PIN Diode current to GNDA

2-3: sink PIN Diode current to PIN_BIAS1

#### JP7 – Alternate Bias Relay Coil Voltage

3-way solder jumper; to reconfigure, pin 1-2 must be cut and pins 2-3 bridged.

1-2: use J2-28 (+12V)

2-3: use J2-32 (TBD)

### Circuit Details

#### Shift Register

The shift register implements the digital control of the bias relays. It consists of three MIC5891 8-bit serial-to-parallel driver chips. All signals driving these chips are 5V TTL level from the Arduino Nano Every on the MBP.

These chips accept a stream of serial data on their Serial Data In pins, and are daisy-chained by connecting Serial Data Out to the next chip’s Serial Data In, making up a single 24-bit shift register. Bits are shifted on the rising edge of SRCLK (MIC5891 CLK pin). When the full 24 bit control word has been clocked, RCLK (MIC5891 STROBE pin) is momentarily brought high, transferring the shift register contents to the output latch. When  is active (low), output buffers with corresponding output latch bits in the “1” state are pulled high. Otherwise they are low (they are not high-Z). Note that is pulled up (inactive) to prevent relay coils from being driven on power-on, before the Arduino has configured the shift register.

The MIC5891 has a logic supply pin (VCC_LOGIC) for its shift register, as well as a supply to be switched to its output pins (VCC_LOAD). The VCC_LOGIC is supplied with +5V, and VCC_LOAD with +12V. Although the bias relays are supposed to contain integrated freewheeling diodes, the MIC5891 outputs also contain transient suppression diodes which appear to be intended to limit the positive voltage on the output pin to a safe level.

The MIC5891 +12V supplies are each decoupled from the +12V supply with a 10Ω resistor, which is shunted by a default-closed copper shunt, and is not populated. If used, under normal operation with a single relay active and drawing 24mA, the series resistor drop would drop 240mV.

#### Bias Relay Circuit

Component references in the explanation below apply to the first circuit in the first group (schematic page 5/30, sheet /Channels/ch0-8/subchan0/).

The bias relay circuits are arranged in 3 groups of 8. Each circuit within a group is controlled by a signal from the group’s MIC5891 shift register chip, called ON0 through ON7 (the fully qualified signal name includes a group hierarchy prefix). These signals switch between GND, and 12V less an output voltage drop of up to ~1V (high). When switched high, the corresponding bias circuit is turned on and the corresponding SiPM channel is activated. Under normal operation, only one channel at a time, among all 24 on the MSB, is intended to be active.

ON0-7 are connected to pulldown resistors, to limit the voltage that will be present, due to leakage current, on an inactive ON signal. Under normal circumstances, i.e. if all relays are present in the circuit, these resistors should not be necessary, as the relay coil will shut any leakage current to ground with a relatively low resistance. These resistors are depopulated.

Bias switching relays are intended to be socketed for ease of troubleshooting and replacement.

The switching in the circuit is accomplished using photoMOS optical relay U5 and 2 Form A (DPST) reed relay K3.

Design goals of this circuit are summarized as follows:

- Connect VB to the SiPM cathode node, when the ON signal is active.

- Isolate VB from the SiPM cathode node when the ON signal is inactive, limiting leakage into SIPM_K to a sub-pA level (exact performance achieved is TBD).

- Provide a 10kΩ source resistance for pulse mode

- Switch current into the RF path downstream of the DC blocking capacitor, to activate downstream PIN diodes, with minimal effect on pulse propagation (i.e. minimal stub length and admittance)

- When active, connect SIPM_K to a measurement signal mode, through a resistor (to minimize effect on pulse propagation, i.e. minimal stub length and admittance)

#### Inactive State

When the ON signal is inactive-low, optical relay U5 and reed relay K3 are both inactive.

To achieve off-state leakage requirements, two stage switching is done using U5 and K3. Typical U5 off-state leakage current is expected to be ~10nA (discuss typical plots from TLP3149 data sheet). This will produce, then, up to 1mV with respect to AGND (BIAS_GUARD) across R11 (100kΩ).

In turn, current driven by this small voltage is blocked by one set of K3 contacts. Based on an off-state reed relay contact resistance of 10GΩ, we could expect a leakage current of 0.1pA.

Note that when the ON signal is low, it is near AGND, so no significant leakage current into the SiPM cathode node is expected.

Note that it is possible that the required leakage level is achievable without using U5. In this case, U5 and R214 can be omitted, and U5 pins 3-4 bridged.

R12 was initially added to provide an additional path to ground, to minimize any residual leakage current passing through K3. However it was found to provide a path for thermoelectrically generated current, and consequently is currently depopulated.

#### Active State – DC path

This circuit has two active-state use cases for IV curve measurement.

First, with a positive VB of up to ~60V, to activate the SiPM in its normal reverse-biased mode, with relatively low current (possibly up to the µA level). This condition is used for IV curve testing as well as pulse counting.

Second, with a negative VB, with voltage limited to 100s of mV, but with currents possibly reaching 100mA levels. This is used only for IV testing of the SiPM in forward bias, to estimate the value of its quenching resistors.

When the ON signal is inactive low, the photomos and reed relays are both active.

Resistors R9 and R10, in series with the current path to the SiPM cathode node, total 20kΩ, and by virtue of the low expected current levels, should result in a small or negligible voltage drop in positive-VB mode.

When VB is made negative, the SiPM is forward-biased and consequently a larger current can be expected to flow in the cathode bias circuit. The voltage drop across R9 and R10 would then be an issue, so D14 is added to bypass these resistors, so a minimal amount of excess voltage is required to achieve forward-biased SiPM voltage needed to measure its forward IV curve. In spite of the D14 voltage drop being manageable, it is still unknown and must be measured (or rather, the resultant SiPM voltage must be measured). To this end, a path is provided, through R190 and the second set of K3 relay contacts, to measurement node VSENSE, which connects through the MBP to a voltmeter.

#### RF Path

This section describes the design requirements for the RF path as regards PIN diode switching.

##### Active State

When the SiPM cathode bias circuit is active, VB is present on the SiPM cathode node SIPM_K.

Low-going pulses originating at the SiPM pass through the TIB and are transferred to a controlled-impedance path consisting of a multi-circuit coaxial ribbon cable assembly, the HSB, and another multi-circuit ribbon cable assembly. The pulses arrive at the MSB on the high-density coaxial connector, where they are transferred to a controlled-impedance microstrip line on the top layer of the board, which is referenced to a ground plane on the first internal layer.

Note that the inter-layer distance and dielectric constant must be considered in calculating the microstrip trace width. These parameters depend on the manufacturer’s stackup, and may be subject to change and should be checked each time the board is ordered, with the trace width adjusted if necessary. The most recent build parameters can be found in text notes in the PCB design file. Any new width can be applied, using a single command, to all traces with netclass ‘zcontrol’.

The pulse passes through capacitor C4, which is located inline with the microstrip.

Current is injected from the active-high ON signal, via R8 and L1, into the RF trace located downstream of DC blocking capacitor C4. This forward-biases and activates two levels of PIN diodes used as RF switches.

The RF path is directly contacted by four other components in the bias relay circuit, all of which are intended to present a high RF impedance to the pulse, and all of which are located close to the microstrip path to minimize stub length. Within the bias relay circuit, these are:

- L1 (DC current source for PIN diode switching)

- D14 (bypass of R9 and R10 for negative VB / forward-biased SiPM use case). A PIN diode is used to minimize capacitance when the PIN diode is not forward biased. Whenever RF performance is required, VB will be positive, and the diode will be essentially zero-biased (by the negligible voltage drop across R9 + R10).

- R10 (VB bias)

- R190 (for measuring VB level)

After the bias relay circuit, these are:

- R49, R90 and R91 (for setting the off-state voltage between the PIN diodes)

- L25 (current drain for PIN diode switching)

The RF part of the multiplexer is implemented with two levels of PIN diode switching. The reason for having the PIN diodes in the first place, is to avoid having the inactive PIN diodes connected, which would act as stubs, smearing out and reflecting the pulses from the active SiPM. Note that the stub lengths back to inactive PIN diodes are minimized, at both the 8:1 and 3:1 junctions.

When the ON signal is active high for a single bias relay circuit, a current is injected via R8 (1kΩ). This current will be approximately 10mA. This will result in a typical on-state resistance of around 0.56Ω, in series with the RF signal.

##### Inactive State

The RF path has no requirements in its inactive state, except for having a minimal effect on the active RF path that it connects to at the junction point between PIN diodes. The relevant parameters here are the stub length from the junction point back to the inactive PIN diodes, and the PIN diode off-state capacitance (see discussion below in this section).

The PIN diode is specified to have a typical zero-bias off-state capacitance of 0.65pF. The first level of inactive PIN diodes will present seven off-state PIN diodes, and the second level will present two off-state PIN diodes, for a total of nine parallel diodes, or typically 5.85pF. This capacitance blocks the signal from traveling back to the inactive SiPMs.

Reverse bias may lower the PIN diode off-state capacitance by 40%. Initially, for the sake of simplicity, this is not done. An explanation of the initial configuration, as well as that required to provide reverse bias, follows.

Zero Bias Configuration

The PIN diode switching system is currently configured for zero off-state bias voltage, with the sink voltage being at ground level. Resistors R49, R90 and R91, as well as R160 and C57, are omitted, and L25 is connected to ground. PIN_BIAS1 and PIN_BIAS2 are not used.

Reverse Bias Configuration

To provide reverse bias, voltages must be provided on PIN_BIAS2 (e.g. 3V) and PIN_BIAS1 (e.g. 6V). In this case, each level of PIN diode will have a reverse bias of 3V. With a 1kΩ source resistance, the activation current will be less (since the sink will be at 6V). It should be possible to set PIN_BIAS1 to 6V simply by choosing an appropriate resistor value to GND (on the MBP).

For this configuration, JP6 must be altered to connect pin 2-3 instead of pin 1-2.

Increased Control Voltage

A reverse bias of ~3V on each diode stage looks feasible if we stick to +12V for the relay activation. However, if we wanted to reduce the off-state capacitance as much as possible (to ~0.4pF), we could increase the ON voltage, possibly swapping out the 12V relays for 24V. This would require a higher voltage from the MSB, to be supplied on J2-32. Doing this, we could provide 8V of reverse bias on each PIN diode, which would reduce the off-state PIN diode capacitance practically to its asymptotic level. Note that the power rating of the series resistor should be double-checked.

It is uncertain at this point whether this effort will be worthwhile; thus it is not done for the initial stage of evaluation.

Insertion Loss and Series Inductance (active PIN diodes)

We want to minimize the insertion loss and series inductance of the two active PIN diodes in the path. The typical insertion loss at 1GHz is given in the data sheet as 0.08dB for a current of 5mA, dropping to 0.065dB at 100mA. We aim for a current of ~5mA. Series Inductance is only specified at the (presumably best-case) 100mA current level, at 0.6nH.

#### Guarding for Low Leakage

We have relatively high voltage nodes (VB, +12V, active SiPM cathode connected to VB) on the surfaces of the board, sometimes in close proximity to nodes into which current leakage would have consequences for the resolution of the current measurement. These sensitive nodes are:

- Bias nodes. We want to minimize leakage into cathode bias nodes of currently inactive SiPMs. Even though leakage current would have to pass through inactive SiPMs to add to the current measurement, we should prevent leakage into the cathodes since we don’t know how easily the current will leak through the inactive SiPMs.

- The IV sense node, which connects to the electrometer. This is the most sensitive node.

Current tracking between PCB traces, exacerbated by thin films of foreign substances that may adhere to the board surfaces, is a possible means for leakage to occur. A well-known technique to prevent such a leakage current path is to place a bare copper trace in its way. This bare copper trace should be driven to a voltage very close to that of the sensitive trace. We employ this method on the MSB.

The MSB has a grounded copper pour zone on both top and bottom layers, so it is not necessary to place traces per se. Instead, we place graphic lines on the solder mask layer, so that the copper is exposed in a continuous barrier between any higher-voltage node and any sensitive node.

The MSB also has internal copper layers. One of these (IN1) is dedicated as a ground plane. IN2 is used for routing. These layers are not subject to external contamination in the same way or degree as the top and bottom layers. However we also have a ground connected copper pour on IN2, which should help to intercept current that might flow horizontally within this layer.

Here are specific leakage paths that are deliberately controlled:

- Within a bias relay circuit, VB to SiPM cathode. To prevent this, there is guarding within each bias relay circuit.

- Between bias relay circuits (one might be active). To prevent this, there is guarding between the circuits.

- SiPM cathode to adjacent SiPM cathode (one might be active). To prevent this, there is guarding between adjacent cathode traces

- Anything, but particularly VB or +12V, into IV output node. To prevent this, there is guarding around the IV output node, from where it comes on the high density connector, through the sense relays, and to the output connector.

R214, U5-1 “ON”: aggressor only against other channels

U5-3 and via, “BIAS”: aggressor against current channel

#### Sense Relay

The sense relays are controlled by IV_ON, a 5V TTL signal.

#### Temperature Readout

Four thermistors are included on the MSB, placed at various locations. Two thermistor connections go to the high density connector, for remote temperature sensing on the HSB. CHECK IF WE HAVE THIS ON HSB OR ELSEWHERE

The goal is to give us an approximate view of temperature gradients in the system.

We use Pt1000 RTDs in an 0603 SMT package. The voltage divider consists of a 4.7kΩ resistor to THERM_ON, with the Pt1000 to GND. All fixed resistors are 0.1% 10ppm/K. This may be overkill but for now it beats doing a more comprehensive error analysis. The ADC voltage is scaled to fit between 0 and 1.024V, which is one of the internal reference settings.

The circuit is active only when the THERM_ON signal is switched high by the Arduino on the MBP, and to avoid self-heating this should be brief (and uniform, with uniform intervals, so any self-heating induced offset doesn’t manifest as noise). The thermistor to be read is selected using THERMSEL0-2, and the reading is done using the Arduino’s 10 bit ADC. For calibration, a nominal 0C calibration signal is available for reading when THERMSEL0-2 is set to 7, produced by substituting a fixed 1kΩ resistor for the Pt1000. GNDA is selected when it is set to THERMSEL is set to 6.

4.7kΩ: Panasonic Electronic Components ERA-3ARB472V, P4.7KBDCT-ND
Pt1000: Yageo Nexensos 32207638, 1759-1011-1-ND

Calculated resolution is better than 0.5K,. Readings should nominally be possible from below 100C to above 50C.

### GNDD connection

GNDD is a separate connection to carry return current for photomos and relays. It is not certain whether this will make any difference. The idea is that this will eliminate voltage drop due to these currents, on GNDA, which may improve IV curve accuracy.

## Mux Board Supports

The Mux Board Supports hold the MSBs securely within the nipple tube. Since the MSBs can hang down from the DSUB-37, it may not be necessary to support them, or it may be feasible to have a something to stabilize them.

## Heat Sink Board (HSB)

The heat sink board consists of straight-through connections between SAMTEC

Nothing other than straight through connections (ideally 50Ω impedance controlled)

Guard traces on board, and skipped circuits on the

Current sense connection to mux

## Tile Interface Board (TIB)

The tile interconnect board adapts between the dual-inline tile connector and the micro-coax connector that connects each SiPM to the heatsink board.

SiPM anodes are all connected together at the TIB, and are capacitively coupled to GND to provide RF signal integrity. Multiple capacitors are likely used, to shorten the RF path from each SiPM anode to ground. For DC (IV curve) operation, these capacitors act as an open circuit, diverting all of the SiPM current to the IV OUT node to be measured by the electrometer.

Rationale for performing this function at the TIB: For best RF signal integrity, the SiPMs should be connected to a coax as closely as possible. But to minimize possible leakage through the coax insulation, we don’t want a voltage difference between the coax conductors, which would be necessary if we individually coupled each cathode to the coax shield after the TIB.

Notes from other section, to be elaborated

Interfaces to tile on one side and HSB on the other side

Both interfaces are on the same side of the board (which will be downwards when mounted)

TBD: 1 or 2 quad interfaces per TIB

50 pin 2mm male on TIB to interface with tile

80 circuit HD coax connector to interface with HSB

Guard traces between K and A row on tile interface connector, and (ideally) between K connections

Every other HD coax connection is skipped to allow for guard trace

Includes capacitor(s) for RF to GND coupling of SiPM pulses

Includes current sense connection to HSB

## Cables

### IV, RF out

One RF out cable is needed per quad (connected to MSB J1).

One IV out cable is needed per quad, if IV outputs are separate (connected to MSB J4)

One IV out cable is needed in total (connected to MBP J5) if IV outputs are combined).

MCX R/A to MCX straight https://www.infinitecables.com/products/rg316-mcx-male-to-mcx-male-right-angle-cable?variant=42809827393777

Infinite Cables doesn’t seem to have R/A MCX to SMA in their catalog, but they would probably make it for us.

These cables go to homemade MCX to SMP adapters, and then into the SMP bulkhead feedthrough.

### MSB to HSB Cable

Description, part number

### HSB to TIB Cable

Description, part number

## SiPM Tile

Each SiPM tile contains 96 SiPMs, organized into four quads of 24. Each SiPM’s A and K connection is separately brought out. Thus, each quad connects to a 50 pin dual-inline connector, of which 48 pins are used for the A and K connections, with the remaining two designated for a ground connection for guarding.

RFI-controlled Ethernet using POE to switch relays

Can use travel router to connect EFOY to wifi (leave it off when not in use)

# Notes, Issues

Multiple bias circuits on at once might over-power R2 (there’s no reason for this to happen, but it could).

Filter VB, ON to prevent noise injection (e.g. via ON, into RF?)

Can this just as well be done on MBP?

Source termination on MBP serial clock.

Handle VSENSE on MBP.

# TDL

☐	Put thermistor pads on HSB and TIB - Tania

☐	Check PCBWay data requirements, include other info such as finish

Last order was HASL / lead

☐	Submit MSB, HSB to PCBWay

☒	Get assembly quote

☒	Optimize bottom side guard aperture placement (and possibly line width)

☐	Make IV_ON work on new backplane design; decide how well we can roll with initial rev BP design (see note on schematic for pin assignment). This means changing IV_ON pins from THERM_OUT1-4 (temporary for rev 0) to D6, D7, D8, D13

This can be done by:

- Removing the jumper on the backplane from pin 36-35 on each DC-37

- Connecting, with a wire, D6, D7, D8 and D13 to pin 35 on J1, J2, J3 and J4 respectively

Lower priority

☐	Block diagram

☐	Cable connection list / discussion

☐	Review / delete / integrate extra stuff in document

Done

☒	Double check zcontrol trace width against old rev, current calculation, PCBWay stackup

☒	Add notes to PCB file regarding TL (CPWG-G) parameters.

☒	Review guarding DR, PCB

(Be informed with how it was done on last rev – maybe not all that relevant)

☒	Check actual PIN diode current on existing board, and match with evenly distributed resistance

It was 1k. In lieu of R5 however, I read 56uH. Maybe it’s nominally 47uH?

☒	Determine inductor value to be used, informed by measuring current board

On built board it measures ~12uH so it’s probably 10uH nominally

☒	Double check BOM (reset status field?)

☒	Add all part numbers

☒	Export to PCB

☒	Put full reference divider (1k / 4300) on temperature channel 7

☒	Fix footprint errors in MSB

☒	Finalize text on board

☒	Go through notes

Check whether 10Ω resistor on MIC5891 +12V is a good idea. It is marked DNP and bypassed by a jumper.

Check VSENSE connection through MBP.

Bias voltages for inactive-state PIN diodes…

Possibly revise R190 value upwards from 10k à done, 68k

Should R5 (after the second PIN diode) really be a resistor, or an inductor? Changed to inductor.

Fix footprint errors?

It would be cool to have a 24V option for the bias relays. Then there would be lots of headroom for reducing the PIN diode capacitance.

Put a jumper option to use a different pin for an alternate voltage

Document the whole PIN diode stuff

Update documentation for alternate voltage

Not being done

☒	Maybe put small positive offset for bottom

# Earlier Text

This is an earlier version of description text. Some of it is to be integrated into section 1.

## Technical Description

This is a signal high level overview of the entire system’s operation in each of its two modes.

### Bias Switching Circuit

The bias control circuit is used to activate the bias on one SiPM at a time, or on one SiPM per MSB at a time.

### IV Mode

### Pulse Mode

In pulse mode, the cathode is biased to a voltage consisting of the breakdown voltage plus a few volts overvoltage. When the SiPM triggers it pulls an amount of charge, which produces a negative-going pulse back along the 50Ω line. This will typically be fed to a CSP and pulse shaper which will stretch the pulse before it is acquired on an ADC. It may also be directly acquired by a faster ADC after some preamplification.

### Components

#### Bias Switch

The function of the bias switch is to take the bias voltage from the voltage source (contained in the B2983A PicoAmmeter) and apply it to one SiPM at a time.

Elements of the IV mux hardware and firmware design are re-used. The following similarities and differences are noted:

- The control logic and firmware are copied from the IV mux design, and the same shift registers are used.

- 24 bits are configured per card, using three cascaded shift registers

- A 16-pin right-angle 2x8 connector, located at the card edge, mates with the motherboard

- Mounting holes to secure the boards, possibly with offsets between them

- A SER_OUT line goes out the connector, but may not be used (it’s feasible to use four streams)

- The switching of the bias voltage uses the same relays and 12V driver chips.

- Layout around the relays will use some guard traces but this may not be as complete as in the original mux design.

#### Motherboard

##### Daughterboard Interface

##### Arduino

- Nano Every, running firmware close to that of the mux

- 4 serial streams

- 4 sense enable outputs on GPIO

##### IV curve current collectors

- Four MCX connectors which come directly from the pulse mux boards

- High isolation relays, in series and possibly also parallel. Will be used only if needed.

- One MCX or triax connector for IV current output to a triax bulhead feedthrough

#### Partitioning

Arduino Nano on a backplane

Four plugin boards. Each board:

- Switches 24 outputs

- Services one 40 circuit microcoax connector

- A total of 96 circuits

#### Interface

Edge connector, 14 pin (keying takes up 2 positions)

SAM11487-ND

| Signals | Count | Description |
| --- | --- | --- |
| VB+,VB- | 2 | Bias voltage |
| GNDA, GNDD, +5V, +12V | 4 | Power |
| SRCLK, RCLK, SRCLR | 3 | Switch control |
| SER_IN (distributed to each board) | 1 | Switch data |
| SER_OUT | 1 | In case we want to daisy-chain |
| Total | 10 |  |

#### Other parts

##### TLP3106A(TP,F

– just 30V

##### G3VM-101HR2(TR05)

Has good data on off-state

##### TLP3109A(TP,F

100V, 1uA = 100Gohm if linear

### Ideas / Notes / Calculations

#### Worst-case IV current loss

This is from leakage through inactive SiPMs.

Assume we have inactive 95 SiPMs in the circuit, and that they present no resistance to the flow of current.

Considering the burden voltage of 20µV and a grounding resistor of 1MΩ. The current is then

95 * 20µV / 1MΩ = 1.9nA

Pretty bad

But, we aren’t populating the 1MΩ resistor…

#### Approaches

            - Group SiPMs, have multiple collection wires, with a few relays

            - Guard using low input bias / low offset drift amp; use switches to allow offset trim; is there a zero-drift IA for trimming?

            - Use guarding using KeySight voltmeter

            - Use SSR with back to back PhotoMOS

#### Calculations for Back-to-Back PhotoMOS

##### Rshunt

It is as small as it can be, for power purposes.

If V=100V and P≤100mW, R≥100kΩ (R=V2/P)

0805, 13k, https://www.digikey.ca/en/products/filter/chip-resistor-surface-mount/52?s=N4IgjCBcoExaBjKAzAhgGwM4FMA0IB7KAbRBgFYwA2ABhHwuqvrIA4rW59zzWB2ACwgAuvgAOAFyggAyhIBOASwB2AcxABffNXggkkNFjyESIAMwBOMzAsjxUyLIUr1WsjVbld%2BwznxFIUjAzAAIAawB5AAsAW0w7EElpAFVlRQkI5ABZbFRMAFd5bE0NDSA

e.g.Panasonic ERJ-P06J133V: 13kΩ, 0805, 1/2W, ok to ~80V

##### Voltage on “off” circuit

Let Ileak=1nA

##### Active Shunt

BSS123WQ-7-F, 100V, SOT-323, Vthmax=2.0V

##### LTC2054/2055

1pA bias current

3uV offset (typ 0.5uV)

##### Use 2-Form-A (back to back) SSR with resistor to GND in middle

Example LCA110STR